## Slide 1: Sunnybrook Summer School: Diffusion MRI
Colleen Bailey, Liam Lawrence, Rachel Chan
August 18, 2025

## Slide 2: Outline
Diffusion Basics
Diffusion Sequences I (diffusion encoding)
  Qualitative: diffusion weighted imaging
  Quantitative: b-values and apparent diffusion coefficient
  Anisotropy: diffusion tensor imaging
Diffusion Sequences II (image encoding)
  Tradeoffs
  Artefacts
Processing Pipeline and Analysis Tools

## Slide 3: Water Motion: A relationship between space and time
“Random walk”
Brownian motion
After time Δt
t=0
Diffusion coefficient
(temperature dependent)
time
Displacement
Δr = √6DΔt
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: What the MRI is measuring here is the effect of molecular diffusion due to Brownian motion. Where molecules will randomly move around from a region of high to low concentration, bouncing around along a trajectory that is sometimes referred to as “random walk” (show trajectory). The rate at which the molecules diffuse is governed by this equation, where the distance diffused is dependent on the diffusion coefficient and the amount of time passed. I won’t say too much more because DWI will be discussed thoroughly in a different lecture.

## Slide 4: A rough idea of scale
Water at 37°C: D ~ 3 µm2/ms
Example Probability Density Functions
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 5: MRI encodes spatial information using gradients The phase and readout gradients encode spatial location Acquire data in k-space Further out in k-space, higher spatial frequencies The diffusion gradients encode spatial location Acquire data in q-space Further out in q-space, more sensitive to small displacements But also dependence on diffusion time
For Imaging
For Diffusion
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 6: A diffusion pulse sequence walkthrough
Read left-to-right as a series of scanner instructions about rf pulses and gradients
Most common diffusion sequence: Pulsed Gradient Spin Echo (PGSE) AKA Stejskal-Tanner diffusion sequence (1965)
To explain how this sequence is sensitive to water motion, consider two situations: without and with motion.
...Image encoding (next section)
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 7: A diffusion pulse sequence walkthrough
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 8: A diffusion pulse sequence walkthrough
90°
Effective diffusion gradient
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 9: Consider two spins in different positions after a 90⁰ pulse
Excitation
90°
position
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 10: Turning on a magnetic field gradient means spins in different positions precess at different frequencies. When the gradient turns off, the spins are dephased.
Excitation
Gradient on
Dephasing
Gradient off
90°
position
dBz
Slow
Fast
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 11: Then wait. During this diffusion time, the spins may move. First consider the case with no motion.
Excitation
Gradient on
Gradient off
Wait
90°
position
dBz
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 12: When gradients are applied in the opposite direction, the left spin now precesses faster than the right. If this gradient is on for the same time, the spins rephase.
Excitation
Gradient on
Gradient off
Wait
Opposing gradient on
Rephasing
Gradient off
90°
position
dBz
Slow
Fast
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 13: If we read out (sum) the signal from all spins in the voxel, they will all be aligned and signal will be high
Excitation
Gradient on
Gradient off
Wait
Opposing gradient on
Rephasing
Gradient off
Acquire signal
90°
position
dBz
Slow
Fast

## Slide 14: Now consider spins that move during the diffusion time
Excitation
Gradient on
Dephasing
Gradient off
Wait
90°
position
dBz
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 15: A dephased spin moves…
Excitation
Gradient on
Gradient off
Wait
90°
dBz
position
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 16: The gradient at the new position is not as large. The spins do not completely rephase.
Excitation
Gradient on
Gradient off
Wait
Opposing gradient on
Partial
rephasing
Gradient off
90°
position
dBz
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 17: Summing all spins at all positions in the voxel, they are no longer in phase and we get a much smaller signal (the signal attenuates more when spins have moved more)
Excitation
Gradient on
Gradient off
Wait
Opposing gradient on
Partial
rephasing
Gradient off
Acquire signal
90°
position
dBz

## Slide 18: When water moves, DWI signal gets lower
How much lower depends on
Scan parameters: gradient amplitude (G), gradient duration (δ), gradient separation (Δ)
Tissue properties
water
cells
Can we go beyond DWI and quantify diffusion?
Where are these parameters on the scanner?
Which should you set and to what values?
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 19: Quantification: b-value and Apparent Diffusion Coefficient (ADC)
The b-value summarizes the scan parameters into one number: b=(γGδ)2(Δ-δ/3)
For Gaussian diffusion: ln(S/S0)=-b*ADC
water
cells
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 20: Quantification: b-value and Apparent Diffusion Coefficient (ADC)
The b-value summarizes the scan parameters into one number: b=(γGδ)2(Δ-δ/3)
For Gaussian diffusion: ln(S/S0)=-b*ADC
Can calculate ADC from 2 b-values
water
cells
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: So you really only need two points to define a line

## Slide 21: Why don't these ADCs for water agree?
water
ADC
DWI
2.1 µm2/ms
1.3 µm2/ms
It's noise!
Always check the DWI *and* the ADC map
SNR limits are one reason diffusion has larger voxels than structural
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: Many scanners automatically generate an ADC map
Could consider a question on SNR. If you go from 1x1x1 mm3 voxel to 2x2x2 mm3, you get 8x the SNR. You would have to scan 64x longer to get the same boost in SNR just from averages. (Lower resolution also allows you to do fewer phase encodes and reduces TE, so maybe this should go later).

## Slide 22: Why don't these ADCs for cells agree?
ADC
DWI
0.6 µm2/ms
0.5 µm2/ms
cells
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 23: Why don't these ADCs for cells agree?
ADC
DWI
0.6 µm2/ms
0.5 µm2/ms
cells
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Water in cells does not undergo Gaussian diffusion at high b!
Even if you see signal in DWI, ADC may not be appropriate at high b (>1000 s/mm2)

## Slide 24: There are hardware and practical limitations
The scanner has a maximum gradient amplitude
TE limits gradient duration and separation
Longer TEs have lower SNR (T2 decay)
A "b-value" may not mean the same thing on different scanners.
Scanner (3 T) | Max Gradient (mT/m)
GE Signa Architect | 44
GE Signa Premier | 80
Philips Ingenia Elition | 45
Philips MR 7700 | 65
Siemens Skyra | 45
Siemens Prisma | 80
Siemens Connectom | 300
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: I wanted to do a little vendor comparison of their max DC gradients and thus the b-value that you would get at a particular diffusion time for their PGSE sequence (although it's complicated if Siemens doesn't even use PGSE by default).
Multi-centre studies are challenging, even with good QA.
Images of scanner GUIs? Not all allow control of all parameters or even let you easily see them

## Slide 25: Example Scanner Interface

## Slide 26: Other Considerations for DWI and ADC
Low b-values may be affected by blood flow in capillaries (modelled by IVIM)
ADC is commonly used for quantification at low b (exact value is scanner-dependent)
Kurtosis or axon diameter mapping can be used for quantification at high b
Slice selection gradients can contribute to diffusion weighting (mainly at low b).
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: (Move the slice select info to artefacts/tips section?)

## Slide 27: Anisotropy and Diffusion Tensor Imaging (DTI)
Data: https://dmri.slicer.org/docs/
Scanners typically have 3 gradient coils
Turn on in various combinations to examine water motion in different directions
Diffusion signal varies with gradient direction (eg. WM fibre orientation)
Unweighted Image
Direction 1
Direction 2
Direction 3
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 28: What is a tensor and what can it tell us?
Can be described by a tensor
Example Probability Density Functions
?
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Requires 6 independent directions (plus unweighted image)
This basic tensor shape doesn't describe all possibilities but works well for aligned axons vs grey matter
Notes: Simple model is the diffusion tensor; a 3x3 matrix that can be conceptualized as an ellipsoid of sorts. You might think that an ellipsoid has three measurements to define it in size, but then you also need three parameters to describe its orientation in space, so overall 6 parameters and you need at least 6 directions, to determine the diffusion tensor, plus an unweighted to determine magnitude.
Now we have a b-matrix rather than a b-value. And a diffusion tensor rather than an ADC. It is 3x3 and symmetric, hence 6 values are needed to populate it.

## Slide 29: DTI – Common Summary Parameters
Crossing fibres
Aligned fibres
Data: https://dmri.slicer.org/docs/
Mean diffusivity
Fractional Anisotoropy
Radial diffusivity
Axial diffusivity
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: FA equation summarizes the amount of anisotropy: 0 is isotropic and 1 would be perfectly aligned. A value of 0.7 is common in aligned WM voxels. This definition of anisotropy is independent of fibre orientation but can compound uncertainties and measurement errors. AD and RD may be less "noisy." Note that none of these tell you the actual direction of the fibres.

## Slide 30: Colour FA: Summarizes FA and Fibre Direction
Data: https://dmri.slicer.org/docs/
MD and FA are rotationally invariant scalar values, which make for easier comparisons between subjects.
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 31: Limitations of DTI
FA is particularly sensitive to noise
Cannot account for multiple fibre orientations within a voxel
  Possible to explore shapes beyond tensor
  Requires more measurements and/or assumptions about the distribution
DTI assumes Gaussian diffusion
  Usually done at one b-value.
  Sets of directions at multiple b-values (shells) are possible but time-consuming
  Quantification options include DKI and NODDI.
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 32: Example Scanner Interface

## Slide 33: Tractography for Fibre Bundles
Tensors in neighbouring voxels can be used to model the likely trajectory of white matter pathways
Prone to false positives. Incorrect tracts need to be edited
There are a variety of algorithms​
  Select a point in a seed region
  Follow the likely local direction a small step size
  Repeat until termination criterion is reached (eg. low FA)
  These steps can be repeated for multiple points in the seed region to generate many streamlines
Christidi et al. 2022. Neurol. Int. 14: 841.
https://pmc.ncbi.nlm.nih.gov/articles/PMC9589952/
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 34: How do we acquire an image?
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: EPI is the reason for diffusion's infamous noisiness.

## Slide 35: How do we acquire an image?
There are many different ways to traverse k-space
The most common method for clinical diffusion imaging uses echo planar imaging (EPI)
It is fast (for MRI) and less sensitive to motion but prone to certain artefacts
...Image encoding
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: EPI is the reason for diffusion's infamous noisiness.

## Slide 36: Echo Planar Imaging (EPI)
Time to traverse k-space ~ 50 ms
Interecho time ~ 0.5 ms
NPE ~ 100
For a spin echo readout, where every phase encode line is done in a separate TR, the acquisition would take 100x as long!
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: There is an interecho time for the gradient echo spacing in the image encoding, and an echo time defining the peak of the spin echo after the 180 pulse.

## Slide 37: Tradeoffs: Diffusion Weighting and Echo Times
Suppose:
Maximum gradient 40 mT/m
100 phase encodes (20 cm FOV, 2 mm resolution)
TE = 60 ms
What is the maximum b-value?
Time to traverse k-space ~ 50 ms
Interecho time ~ 0.5 ms
NPE ~ 100
b=(γGδ)2(Δ-δ/3)
b=(2.68e8*0.04*0.005)2*(0.03-0.005/3)
25 ms
TE/2=30 ms
δ=5 ms
Δ=30 ms
b=8.11e7 s/m2
b=81 s/mm2
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 38: Exercise: Help me get more diffusion weighting
Option 1: Increase TE to 100 ms
Time to traverse k-space ~ 50 ms
Interecho time ~ 0.5 ms
NPE ~ 100
25 ms
TE/2=50 ms
δ=25 ms
Δ=50 ms
b=(γGδ)2(Δ-δ/3)
b=2982 s/mm2
Why would we not want to do this?
T2 decay and lower SNR
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 39: Exercise: Help me get more diffusion weighting
Option 2: Lower resolution 4 mm
Time to traverse k-space ~ 25 ms
Interecho time ~ 0.5 ms
NPE ~ 50
12.5 ms
TE/2=30 ms
δ=17.5 ms
Δ=30 ms
b=(γGδ)2(Δ-δ/3)
b=847 s/mm2
What are the advantages and disadvantages of this?
Lower spatial resolution
4x as much SNR
Would need to scan 16x a long for equivalent SNR gain
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 40: Exercise: Help me get more diffusion weighting
Option 3: Turn on all 3 gradients
Time to traverse k-space ~ 50 ms
Interecho time ~ 0.5 ms
NPE ~ 100
25 ms
TE/2=30 ms
δ=5 ms
Δ=30 ms
b=(γGδ)2(Δ-δ/3)
b=243 s/mm2
Why would we not want to do this?
This only works in 4 directions
Gall=sqrt(3)*0.04
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 41: Exercise: Help me get more diffusion weighting
Time to traverse k-space ~ 30 ms
Interecho time ~ 0.5 ms
NPE ~ 60
Option 4: Partial Fourier Encoding of 0.6
5 ms
TE/2=30 ms
δ=25 ms
Δ=30 ms
b=(γGδ)2(Δ-δ/3)
b=1550 s/mm2
Parallel imaging is another option to reduce the number of phase encodes:
Multiple receive coils
SENSE or GRAPPA
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Notes: Is it worth having another slide for parallel imaging just to show that you're covering all of k-space, but only every other line (but with two coils and then you reconstruct)

## Slide 42: Exercise: Help me get more diffusion weighting
Template
Time to traverse k-space ~ 50 ms
Interecho time ~ 0.5 ms
NPE ~ 100
25 ms
TE/2=30 ms
δ=5 ms
Δ=30 ms
b=(γGδ)2(Δ-δ/3)
b=2982 s/mm2
Why would we not want to do this?
T2 decay and lower SNR

## Slide 43: A note about slices
Examples above are for one slice
EPI data is often acquired "multi-slice" not "3D"
  Slice selection + 2D k-space
  Repeat the slice-selective excitation for different slices during TR
  More slices require longer TR, so there is some increase in imaging time
It is possible to acquire 3D k-space and do a 3D Fourier transform
  Requires long echo times unless there are few phase encodes
  Has advantages for resolution
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 44: EPI Artefacts: Susceptibility-based distortion
Off-resonance distribution at 3T
Images: Fa-Hsuan Lin
TSE
EPI
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 45: EPI Artefacts: Susceptibility-based distortion
The bandwidth per pixel in the readout direction is: 1/0.0005 = 2000 Hz/px
In the phase direction, bandwidth per pixel is: 1/0.05 = 20 Hz/px
Susceptibility that varies the B0 field by 100 Hz creates a shift:
  100/2000 = 0.05 pixels in readout
  100/20 = 5 pixels in phase
Appears as stretch or compression in phase encode direction
Tread~0.5 ms
Tphase~50 ms
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 46: EPI Artefacts: Susceptibility-based distortion
Tread~0.5 ms
Tphase~50 ms
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
Fixes
Reduce susceptibility
  Shimming
  Thinner slices
Correct for susceptibility
  Acquire a B0 map and correct in post-processing or use to shim
  Acquire an image with "reverse blip"
Reduce # phase encodes
  Partial Fourier
  Parallel imaging
  Choose the "right" phase encode direction

## Slide 47: Chemical Shift Artefact
Fat is shifted 3.5 ppm from water, about 450 Hz at 3 T
For readout bandwidth of 2000 Hz/px, this is 0.2 pixels
For phase bandwidth of 20 Hz/px, this is 22 pixels
Fat signal attenuates less than water, so artefact may be more obvious in DWI than unweighted image.
Use scanner fat saturation
Chemical shift
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 48: Eddy current distortions
Turning on and off gradients induces eddy currents, changes effective gradient
Two problems:
  Diffusion gradients aren't what we plan (altered b-value)
  EPI readout will not be as expected, creating artefacts
This effect will be worse for stronger diffusion gradients
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 49: Eddy Current Distortions - Readout
Shear
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 50: Eddy Current Distortions - Phase
Stretch
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 51: Eddy Current Distortions - Slice
Translation
These are all approximately affine transformations.
Can be corrected through affine registration.
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 52: EPI ghosts
k-space acquisition alternates direction for even and odd lines
Small timing differences produce phase differences
These appear as ghosts in image space
Always in the phase direction, a distance of N/2 apart
Many scanners have automatic correction based on pre-scan or navigators
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 53: Potential Eddy Current Solutions
Worse at high gradients
Use a "self-compensating" sequence like TRSE
  Longer TE for a given b-value
Correction acquired during pre-scan
Distortions can often be fixed with affine registration
  Can be challenging as contrast changes
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 54: Multi-Shot Imaging and Motion
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline
EPI phase encodes can be split into separate TRs (multi-shot EPI)
Shots may be seconds apart vs 50 ms for single shot
Motion between shots causes phase errors
  Brain pulsations from heart
  Cardiac gating adds scan time
  Acquire a navigator for phase correction

## Slide 55: Multi-Shot Imaging and Motion
Multi-shot EPI with interleaved phases may have gaps in k-space after correction
Could leave residual ghosting
Readout-segmented EPI is an alternative (eg. RESOLVE)
Overlap segments so there are no gaps after correction
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline

## Slide 56: Image Processing Steps and Analysis Tools
Look at the data (SNR, artefacts)
Susceptibility correction with B0 map or reverse blip
Motion correction
  b-vector correction
Eddy current correction
Segment brain
Apply model (eg. DTI)
Calculate metrics (MD, FA)
Compare subjects
Many scanners automatically do eddy current correction, calculate ADC and FA
Software often has its own tutorial
  FSL (Oxford)
  DTIStudio (Johns Hopkins)
  Camino (UCL)
  3DSlicer with SlicerDMRI extension (NIH), Tutorial.
Diffusion Basics
Sequences I: Diffusion Encoding
b-value and ADC
Sequences II: Image Encoding
Anisotropy and DTI
Sequences II: Tradeoffs
Sequences II: Artefacts
Processing Pipeline